## Instructions
Overview
This workbook contains a template for the calibration of Mass Flow Controlled (MFC) for high volume particulate sampler for PM10 or TSP.
The calculations in this worksheet are based on EPA Method IO-2.11 and are intended to be compatible with manufacturer recommended methods.
Instructions
Fill in the Orifice Qa Calibration worksheet with the calibration data points for the orifice being used. Alternatively, the Orifice ID, Calibration Date, Slope and Intercpet may be entered manually in the appropriate fields of the Cal Form.
Protection
The worksheets are protected and only cetain cells have been unlocked. Some unlocked cells may contain formulas or references which may be desirable to overwrite with manual entries.
If the sheets need to be edited beyond the unlocked fields, the protection may be turned off in the File menu or Review tab depending on version of Excel. Protection is not password locked.

## Cal Form
Hi-Vol (PM10 or TSP)
Particulate Sampler Calibration Verification
Mass Flow Controller
| Site |  |  |  |  | Calibration Orifice
Location:
AIRS ID | [5 empty] | Serial:
| POC: |  |  |  |  | Cal Date:
Date: | [5 empty] | Qa Slope:
Sampler Serial #: | [5 empty] | Qa Int.:
Verified By:
Pre Verification/Maintenance
Calibration Data Results |  |  | Calibration Date: |  | Orifice Used: |  | Orifice Cal Date:
|  | Point/Plate 1 | Point/Plate 2 | Point/Plate 3 | Point/Plate 4 | Point/Plate 5 | Slope:
Qa (m3/min): | [6 empty] | Intercept:
(Ic) Corrected Dickson | [6 empty] | Corr Coef:
Check Standard SN / Cert: | [5 empty] | Instrument:
Ta (deg F): |  |  |  |  | Setpoint Flow:
Ta (deg C): |  |  |  |  | Last Average Dickson Reading:
Ta (deg K): |  |  |  |  | Reason for Check:
Pa (mm Hg):
Plate/Point |  | Orifice | Orifice Acutal | Dickson | Qa | Flow | PM 10
Number |  | inches H2O | m3/min | Recorder | Calculated | % Difference | Δ a%:
Point 1 (PM10 or TSP) | [6 empty] | Δ b%:
Point 2 (TSP) | [6 empty] | Δ c%:
Point 3 (TSP)
Point 4 (TSP) | [6 empty] | TSP
Point 5 (TSP) | [6 empty] | Overall % Diff:
Calculations
Calibrator Flow (Qa) = 1 / Slope * (SQRT(Orifice * (Ta / Pa)) - Intercept)
Orifice is orifice pressure in inches of Water Ta is ambient temperature in degrees Kelvin Pa is barometric pressure in mm Hg
For subsequent calcuations of sampler flow: 1/m((I)(Sqrt(Tav/Pav))-b)
* Monthly check should be done to see if measured flow rate is within +/- 7% of calculated flow and +/- 10% of setpoint (PM10 only)
Post Verification/Maintenance
Ta (deg F): |  |  |  |  | Reason for Verification:
Ta (deg C): |  |  |  |  | Pa (mm Hg):
Ta (deg K):
Plate/Point | Orifice | Qa | (I) Dickson | (Ic) Dickson |  | Regression Results
Number | inches H2O | m3/min | Chart | Corrected |  | Slope:
1 | [5 empty] | Intercept:
2 | [5 empty] | Corr Coef:
3
4 | [5 empty] | PM 10 Seasonal Setpoints (Dickson)
5 | [5 empty] | DEC/JAN/FEB
MAR/APR/MAY-N8
Comments: | [5 empty] | MAR/APR/MAY-S8
JUN/JULY/AUG/SEP
OCT/NOV
TSP Set Point:
Verification of operational flow rate
Point |  | Orifice | Qa | Dickson | Qa | Flow | Δ a%:
Number |  | inches H2O | m3/min | Recorder | Calculated | % Difference | Δ b%:
Point 1 (PM10 or TSP) | [6 empty] | Δ c%:
Δ a% = Orifice vs Setpoint <10% (PM10 only)
Δ b% = Dickson Calculated vs Setpoint <10% (PM10 only)
Δ c% = Orifice vs Dickson Calculated <7%

## Cal Form No Equations (hidden)
Hi-Vol (TSP or PM10)
Particulate Sampler Calibration
Volumetric Flow Controller
|  | Site |  |  | Calibration Orifice
Location: |  |  |  |  | Serial:
Airs ID |  |  |  |  | Cal Date:
Date: |  |  |  |  | Qa Slope:
Sampler Serial #: |  |  |  |  | Qa Int.:
Calibrated By:
Temp (deg F): |  |  |  | Pf (Clean Filter):
Ta (deg C): |  |  |  | Pa (mm Hg):
Ta (deg K): |  |  |  | Reason for Cal:
Pre-Maintenance
Point | Orifice | Qa | Stagnation | Pf |  | Look Up | % of
Number | "H2O | m3/min | "H2O | mm Hg | Po/Pa | m3/min | Diff
1
2
3
4
5
Calculations
Calibrator Flow (Qa) = 1/Slope*(SQRT(H20*(Ta/Pa))-Intercept)
Pf in mm Hg = in H2O * 25.4/13.61
Pressure Ratio (Po/Pa) = 1-Pf/Pa
% Difference = (Look Up Flow-Calibrator Flow)/Calibrator Flow*100
Post-Maintenance
Point | Orifice | Qa | Stagnation | Pf |  | Look Up | % of
Number | "H2O | m3/min | "H2O | mm Hg | Po/Pa | m3/min | Diff
1
2
3
4
5
Initial Certification check 5 points at, above and below Pf of a clean filter
Field Flow Check: Check 1 ponit at the Pf of a clean filter

## Orifice Qa Calibration
Orifice Calibration Form for terms of Qa |  |  |  |  | Fill in the gray squares
Date
Orifice ID
Ta (C)
Ta (K) | 273
Pa (mm Hg)
Point | Known Volume Vm (M3) | Time (minutes) | Time (seconds) | Time minutes - decimal | Roots Press. P (mm Hg) | Orif. Press Drop H (inches H2O)
1 |  |  |  | 0
2 |  |  |  | 0
3 |  |  |  | 0
4 |  |  |  | 0
5 |  |  |  | 0
6 |  |  |  | 0
7 |  |  |  | 0
8 |  |  |  | 0
9 |  |  |  | 0
Va | Qa | y axis
#DIV/0! | #DIV/0! | #DIV/0!
#DIV/0! | #DIV/0! | #DIV/0!
#DIV/0! | #DIV/0! | #DIV/0!
#DIV/0! | #DIV/0! | #DIV/0!
#DIV/0! | #DIV/0! | #DIV/0!
#DIV/0! | #DIV/0! | #DIV/0!
#DIV/0! | #DIV/0! | #DIV/0!
#DIV/0! | #DIV/0! | #DIV/0!
#DIV/0! | #DIV/0! | #DIV/0!
Use Qa slope and intercept for Volumetircally controlled samplers
Qa slope (m) | #DIV/0!
Qa intercept (b) | #DIV/0!
coefficient (r ) | #DIV/0!

## Lab Calibration (Dummy) (hidden)
Dummy Calibration for lab entry |  | Fill in the gray squares
Site Name
AIRS ID# (55-###-####)
Calibration Date
North or South of Hwy 8
Sampler S/N
POC number
TSP or PM10
Dickson Set Point
Filter Pressure (clean filter) in H2O
Seasonal Barometric Pressure
Orifice #
| Po/Pa | Temperature | Qactual | Qtheoretical
Dec, Jan, Feb | #DIV/0! | -8.3 |  | 0
Mar, Apr, May | #DIV/0! |  |  | #VALUE!
Jun, Jul, Aug, Sep | #DIV/0! | 18.6 |  | 0
Oct, Nov | #DIV/0! | 5 |  | 0
Average |  |  | #DIV/0! | #VALUE!
Point | Dickson Chart | Qtheoretical Flow | Orifice
1 | 0 | #VALUE!
2 | 0 | #VALUE!
3 | 0 | #VALUE!
4 | 0 | #VALUE!
5 | 0 | #VALUE!
Note on B10: Jason Treutel: 44 for TSP samplers 36 for PM10 samplers
Note on D15: Jason Treutel: Lookup for the sampler specific venturi S/N
Note on C17: Jason Treutel: 6.8 degrees south of Hwy 8 3.1 degrees north of Hwy 8
Note on D24: Jason Treutel: Lookup from Orifice Calibration Report - to the nearest 0.10 pressure differential